## Slide 1: Robotics 1
Robot components:
Proprioceptive sensors
1

## Slide 2: Properties of measurement systems - 1
2
accuracy
agreement of measured values with a given reference standard (e.g., ideal characteristics)
repeatability
capability of reproducing as output similar measured values over consecutive measurements of the same constant input quantity
stability
capability of keeping the same measuring characteristics over time/temperature (similar to accuracy, but in the long run)

## Slide 3: Accuracy and Repeatability
low accuracy low repeatability
*
***
high accuracy high repeatability
Robotics 1
3
low accuracy high repeatability
*
***
*
average value
*
target value (ground truth)
*
*
*
*
measurements
better components!
calibration!

## Slide 4: Accuracy and Repeatability
in robotics
accuracy is how close a robot can come to a given point in its workspace
  depends on machining accuracy in construction/assembly of the robot, flexibility effects of the links, gear backlash, payload changes, round-off errors in control computations, ...
  can be improved by (kinematic) calibration
repeatability is how close a robot can return to a previously taught point
depends only the robot controller/measurement resolution
both may vary in different areas of the robot workspace
  standard ISO 9283 defines conditions for assessing robot performance
  limited to static situations (recently, interest also in dynamic motion)
  robot manufacturers usually provide only data on “repeatability”
simple test on repeatability of a Fanuc ArcMate100i robot (1.3 m reach)
4
video

## Slide 5: Properties of measurement systems - 2
5
linearity error
maximum deviation of the measured output from the straight line that best fits the real characteristics
  as % of the output (measurement) range
offset error
value of the measured output for zero input
  sometimes not zero after an operation cycle, due to hysteresis
resolution error
maximum variation of the input quantity producing no variation of the measured output
  in absolute value or in % of the input range

## Slide 6: Linearity, Offset, Resolution
X
input range
output range
Y0 = offset
Xm
Y
Ym
X = resolution
Y Ym
= linearity error
Robotics 1
6

## Slide 7: Sensor measurements some non-idealities
Asymmetry
Bias
Dead zone
Nonlinearity
Scaling factor
Quantization

## Slide 8: Classes of sensors for robots
proprioceptive sensors measure the internal state of the robot (position and velocity of joints, but also torque at joints or acceleration of links)
  kinematic calibration, identification of dynamic parameters, control
exteroceptive sensors measure/characterize robot interaction with the environment, enhancing its autonomy
(forces/torques, proximity, vision, but also sensors for sound, smoke, humidity, …)
  control of interaction with the environment, obstacle avoidance in the workspace, presence of objects to be grasped, …
  mobile-base robots: localization in a map, navigation in unknown environments, …

## Slide 9: Position sensors
provide an electrical signal proportional to the displacement (linear or angular) of a mechanical part with respect to a reference position
linear displacements: potentiometers, linear variable- differential transformers (LVDT), inductosyns
angular displacements: potentiometers, resolvers, syncros (all analog devices with A/D conversion), optical encoders (digital), Hall sensors, ...
the most used in robotics, since also linear displacements are obtained through rotating motors and suitable transmissions

## Slide 10: Absolute encoders
rotating optical disk, with alternate transparent and opaque sectors on multiple concentric tracks
(infrared) light beams are emitted by leds and sensed by photo-receivers
light pulses are converted into electrical pulses, electronically processed and transmitted in output
o
resolution = 360 / 2
Nt
digital encoding of absolute position
when the optical disk is rotating fast, the use of binary coding may lead to (large) reading errors, in correspondence to multiple transitions of bits
Photo-receiver
Photo-emitter
Optical disc
Photo-receiver
multiple bit “race” zone
Nt = # tracks = # bits (min 12 in robotics)

## Slide 11: Absolute encoding
GRAY CODE
BINARY CODE
adjacent codes differ by just one bit
XOR
8-bit Gray-coded absolute encoder
binary coding
optical disks with 2 bits
Gray coding
DECIMAL BINARY GRAY

## Slide 12: Use of absolute encoders
ready to measure at start (no “homing”)
two modes for permanent operation
  when switching off the drive, position parameters are saved on a flash memory (and brakes activated)
  battery for the absolute encoder is always active, and measures position even when the drive is off
  data memory > 20 years
single-turn or multi-turn versions, e.g.
  13-bit single-turn has 213 = 8192 steps per revolution (resolution = 0.044o)
  29-bit multi-turn has 8192 steps/revolution
+ counts up to 216 = 65536 revolutions
aluminum case with possible interface to field bus systems (e.g., CANopen or PROFIBUS)
typical supply 5/28V DC @1.2 W
hollow shaft
round flange
13-bit absolute encoder opened: Gray-coded disk and electronics

## Slide 13: Incremental encoders
optical rotating disk with three tracks, alternating
transparent and opaque areas: measures incremental angular displacements by counting trains of Ne pulses (“counts”) per turn (Ne = 100÷5000)
the two A and B tracks (channels) are in quadrature (phase shift of 90o electrical), allowing to detect the direction of rotation
a third track Z is used to define the “0” reference position, with a reset of the counter (needs “homing” at start)
some encoders provide as output also the three phases needed for the switching circuit of brushless motors
The three tracks on an optical disk (here Ne = 6)
Photo-receiver
Photo-emitter
Optical disk

## Slide 14: Incremental encoders
two (cheap) incremental encoders inside a mouse
a OMRON incremental encoder with 2000 pulses/turn
diameter ∅ 40 mm mass m ≈ 100 g inertia J = 1⋅10-6 kg m2

## Slide 15: Signal processing
“fractions of a cycle” of each pulse train are measured in “electrical degrees”
1o electrical = 1o mechanical/Ne 360o mechanical = 1 turn
signals are fed in a digital counter, with a D-type flip-flop to sense direction + reset
to improve resolution (4×), the leading and trailing edges of signals A and B are used
the sequence of pulses C will clock now the counter (increments or decrements)
C
signals on channels
commutation phases for AC motors
electrical degrees
Ch A
Ch B
D Q
cl
0 = CCW
1 = CW
Ch A
Ch B
D Q
cl
CCW/CW
U/D
cl reset
⇒
Ch Z
n bits
Ch A Ch B
C
C
a
R
a
time delay
 ≈ 45° electrical

## Slide 16: Count multiplication example of quadrature detection
an incremental encoder with Ne = 2000 (electrical) cycles provides a count of N = 8000 pulses/turn after electronic multiplication
its final resolution is (mechanical) 360o/8000 = .045o (= 0o 2’ 42’’)
needs a 13-bit counter to cover a full turn without reset (213 = 8192)

## Slide 17: Quadrature detection in incremental encoders a more complete implementation
it is assumed that an oversampling clock "clk” (e.g., as provided by a FPGA) is available, which is faster than the two quadrature signals A and B
the digital count output will have a resolution multiplied by 4
NOTE: since in practice A and B signals may
not be synchronous to the clock signal, two extra D flip-flops per input should be used
to avoid meta-stable states in the counters
90° electrical
XOR gates
D (delay) flip-flops

## Slide 18: Accuracy in incremental encoders
division error: maximum displacement between two consecutive leading/trailing edges, typically within max
± 25o electrical
the phase shift of the two channels, nominally equal to 90o electrical, is typically within max ± 35o electrical (quadrature error)
...apart from quantization errors
electrical degrees

## Slide 19: Indirect measure of velocity
convolution filtering is needed because of noise and position quantization
  use of non-causal filters (e.g., Savitzky-Golay) helps, but introduces delays
Kalman filter for on line state estimation (optimal, assuming Gaussian noise)
𝑞
Robotics 1
19
animation of Savitzky-Golay filter with cubic polynomials
numerical differentiation of digital measures of position
  to be realized on line with Backward Differentiation Formulas (BDFs)
problems of peaks at low speed
𝑞
̇ 𝑞̇&
𝑇
1
1-step BDF (Euler): 𝑞̇ = 𝑞̇ 𝑘𝑇 = 𝑞 − 𝑞
' ' '-.
directly from incremental encoder
'
⟺ 𝑞̇ =
01
2
3
⟸
1 25
4-step BDF: 𝑞̇ = 𝑞
' '
− 4𝑞
𝑇 12 3 4
+ 3𝑞 − 4 𝑞 + 1 𝑞
'-. '-: '-; '-<

## Slide 20: Kinematic Kalman Filter for velocity estimation
actual state
noisy position measure (encoder output)
unmeasured velocity
= sampling time
zero mean Gaussian noises with (co)variances (a matrix) and
motion and sensing discrete-time model for estimation
[mm]
[mm/s]
design a (linear) Kalman filter providing an estimate
of the model state
using the optimal Kalman gain
(a priori) prediction correction (based on the measured output)
position measure and its filtered version numerical velocity and its filtered estimate

## Slide 21: Velocity sensor: Tachometer always mounted on the (electrical) motor axis
N
S

permanent magnet
coil with area 

principle of operation (single coil)


V
B = cost
amplitude V  
⇒ to reduce ripples, use m coils rotated regularly by 180o/m
magnetic field:
flux through the coil is (B) = |B| cos   |B| cos t
V = - d/dt = |B|  sin t

## Slide 22: DC tachometer an example
Robotics 1
22
Servo-Tek Tach Generator (B series)
bi-directional
output voltage 11÷24 V @1000 RPM
low ripple: < 3% peak-to-peak of DC value (with 72 KHz filter)
weight = 113 g, diameter = 2.9 cm
linearity error < 0.1% (at any speed)
stability 0.1% (w.r.t. temperature)
1.75 mNm (as a load)

## Slide 23: Accelerometers
measure of linear acceleration based on inertial forces (no “touch”)
  units: [m/s²] or gravitational acceleration [g] (non-SI unit: 1g ≈ 9.81 m/s²)
different principles for converting mechanical motion in an electrical signal
  piezoelectric: piezoceramics (PZT) or crystals (quartz), better linearity & stability, wide dynamic range up to high frequencies, no moving parts, no power needed
  piezoresistive: for high-shocks, measures also static acceleration (g0), needs supply
  capacitive: silicon micro-machined sensing element, superior in static to low frequency range, can be operated in servo mode, cheap but limited resolution
  modern solution: small MEMS (Micro Electro-Mechanical Systems)
multiple applications: from vibration analysis to long range navigation
animation of measurement principle in a piezoelectric accelerometer

## Slide 24: Operation principle
seismic accelerometer
by Laplace
transform
sensor
seismic mass
accelerometer
case
surface of moving (or vibrating) body
accelerometer plate
measure
piezoelectric element
base
seismic mass
spring
output
© Brüel & Kjaer
𝑎
>
𝑑:𝑧
𝑎 = 𝑑𝑡:
𝛿
𝑀𝑎 = 𝑀𝛿̈ + 𝐷𝛿̇ + 𝐾𝛿
𝑎> = 𝐾D 𝛿
𝐴>(𝑠) = 𝐾
𝐴(𝑠)
D 𝑀𝑠: + 𝐷𝑠 + 𝐾
LM
𝑀
=
NOP R⁄S NP L⁄S

## Slide 25: Frequency characteristics
of a piezoelectric accelerometer
r
useful bandwidth 𝐵] ≈ 0.3𝜔X
for 10% error
- 3 db = 0.7071
B-3
bandwidth
Bu
resonance frequency
damping factor
low
low frequency limit
of piezoelectric use (≤ 1 Hz)
model valid above this
frequency
1.1
0 dB = 1
𝐹 𝑠 = UV(N) =
U(N)
LM
NOP R⁄S NP L⁄S
X
𝜔 =
𝐾Y
𝑀
𝜁 = 𝐷
2
1Y
𝐾𝑀
𝐹(𝑗𝜔)

## Slide 26: MEMS accelerometers
very simple MEMS (a cantilever beam with a test mass, with damping from the residual gas sealed in the device), single- or tri-axial, very small and light
cross-couplings among acceleration sensing directions should be limited ≤ 3%
ADXL335 3-axis, small, low power, ±3g, with signal conditioned voltage outputs

## Slide 27: Mounting accelerometers on robots
3-axial MEMS accelerometer on the forearm
of a KUKA KR15/2
[DLR/Sapienza, 2007]
3-axial capacitive accelerometer on end-effector tool of an ABB robot (Crossbow Technology: 2g range,
1V/g output, 0-50 Hz, ±2° align error) [Linköping, 2012]
Bosch BMA 150 3-axial accelerometers integrated in two larger Tactile Modules on the links of a Bioloid humanoid left arm [TUM, 2011]